## Appendix 3
Detection Method |  |  | Papers mentioned in/30 | Technology readiness level (TRL) for wastewater | Cost NR = Not reported | Pathogens detection | Sample | Sensitivity | Specificity | LOD (Limit of Detection) | References
Culture-Based Methods | Culture alone |  | 9 | 8 to 9 | expensive | C. coli; C. jejuni; C. lari |  | 52.6% | Very high | NR | Zhang et al., 2023
| Culture with biochemical identification | Biochemical tests: ADA-V | 1 | 8 to 9 | NR | Aeromonas spp. | Natural reservoir | 81.3 | NR | NR | Milligan et al., 2023
|  | Biochemical tests: Pseudomonas CN |  |  |  | P. aeruginosa | Natural reservoir | 100 | NR | NR | Milligan et al., 2023
|  | Biochemical tests: Cetrimide agar |  |  |  | P. aeruginosa | hospital waste water effluent | 93.1 | NR | NR | Milligan et al., 2023
|  | API 20 NE; isolation media: MDR-CA |  |  | NR | Acinetobacter spp. | Riverwater and sediment | 31.3 [Confirmation rate (%)] | NR | NR | Milligan et al., 2023
|  | Biolog MicroLog; isolation media: AIA |  |  |  | Aeromonas spp. | Urban playa lake | 100 | NR | NR | Milligan et al., 2023
| Culture and nucleic based methods | Culture + PCR | 1 | 7 |  | C. coli; C. jejuni | wastewater | NR | NR | NR | Zhang et al., 2023
|  | Sequenced 16S rRNA gene; isolation media: Cetrimide agar |  |  | NR | P. aeruginosa | natural hospital wastewater effluent | 86.2 | NR | NR | Milligan et al., 2023
|  | Sequenced gyrB, radA; isolation media: RYAN |  |  |  | Aeromonas spp. | wastewater effluent | 46 | NR | NR | Milligan et al., 2023
|  | Sequenced gyrB, radA; isolation media: RYAN |  |  |  | Aeromonas spp. | estuary | 33 | NR | NR | Milligan et al., 2023
|  | Sequenced gyrB; isolation media: AIA |  |  | NR | Aeromonas spp. | Urban playa lake | 100 | NR | NR | Milligan et al., 2023
|  | PCR of 16S rRNA gene variable regions 2 and 8; isolation media: Pseudomonas CN |  |  |  | P. aeruginosa | natural reservoir | 100 | NR | NR | Milligan et al., 2023
|  | PCR of gyrB; isolation media: PIA |  |  |  | P. aeruginosa | surface water and wastewater | 73.3 | NR | NR | Milligan et al., 2023
| Multiple Tube Fermentation (MTF) |  | 1 | 7 | low cost | E coli. | wastewater | NR | NR | 1 × 10^3 CFU/L | Chen et al., 2024
| Culture and latex agglutination |  | 1 | 7 | NR | Staphylococcus aureus | wastewater | NR | NR | NR | Chen et al., 2024
| cell culture method, neutralization test, and immunofluorescence method |  | 1 | 6 | NR |  |  | NR | NR | NR | Chen et al., 2024
Nucleic Acid-Based Methods (Molecular Techniques) | Polymerase Chain Reaction (PCR) | Conventional PCR | 30 | 8 to 9 | costly | enteroviruses, HAV, E. coli, Cryptosporidium, Giardia, etc. | wastewater | Very high | Very high | NR | Singh et al., 2024
[6 empty] | C. coli; C. jejuni | wastewater | NR | NR | NR | Zhang et al., 2023
[6 empty] | Cryptosporidium | wastewater | NR | NR | NR | Zahedi et al., 2021
|  | Quantitative PCR (qPCR) | 30 | 8 to 9 | majority of thermal-cyclers required for qPCR start at $20,000 | E. coli, Klebsiella pneumoniae, Clostridium perfringens and Enterococcus faecalis | wastewater | Excellent (0.45 cysts per reaction for the detection of G. lamblia and Giardia ardeae in wastewater samples) | Excellent | NR | Singh et al., 2024
[6 empty] | Campylobacter spp. (Campylobacter jejuni, Campylobacter coli , Campylobacter lari etc) | wastewater | NR | NR | NR | Zhang et al., 2023
[6 empty] | Monkeypox | wastewater | 10 % (LOD defined by minimum Ct with 95% likelihood of detection); 64 % (LOD defined by any qPCR fluorescence) | 98 % (LOD defined by minimum Ct with 95% likelihood of detection); 90 % (LOD defined by any qPCR fluorescence) | NR | Julian et al., 2024
[6 empty] | Cryptosporidium and Giardia, K. pneumoniae, E. coli, E. faecalis, and C. perfringens |  | NR | NR | NR
[6 empty] | P. aeruginosa | natural river | 93 | Very high | NR | Milligan et al., 2023
[6 empty] | human adenovirus |  | NR | NR | NR
|  | Digital PCR (dPCR) | 5 | 8 | Higher than other PCR | SARS-CoV 2 variants, quantify faecal markers | wastewater | Very high | Very high | NR | Singh et al., 2024
|  | Nested PCR | 1 | 8 | costly | SARS-CoV-2 | wastewater | Very high | Very high | NR | Sangkham et al., 2021
|  | Reverse Transcription PCR (RT-PCR) | 7 | 8 to 9 | costly | Adenovirus | medical wastewater | Very high | Very high | LOQ: 1.5 × 10^1 | Chen et al., 2024
[5 empty] | NR | HAV | medical wastewater | NR | NR | 95: 8.02 × 10^2~ 2.80 × 10^3 copies/mL | Chen et al., 2024
[5 empty] | NR | Norovirus | medical wastewater | NR | NR | 8.02 × 10^2~ 2.80 × 10^3 copies/mL | Chen et al., 2024
[5 empty] | NR | salmonella | medical wastewater | NR | NR | 7.5 × 10^−2 CFU/L | Chen et al., 2024
[5 empty] | NR | SAR-CoV-2 | medical wastewater | NR | NR | 5 × 10^1 copies/mL | Chen et al., 2025
[5 empty] | NR | Aspergillus fumigatus | medical wastewater | NR | NR | 40 fg; LOQ: 400 fg | Chen et al., 2026
|  | Reverse Transcription qPCR (RT-qPCR) | 7 | 8 to 9 | NR | SARS-CoV-2 | swab/sputum | 100 | 99.26/ 99.83 | 1000 TUa/mL | Mahmoudi et al., 2022
[5 empty] | NR | C. jejuni | wastewater | NR | NR | NR | Zhang et al., 2023
[5 empty] | NR | SAR-CoV-2 | wastewater | most reliable | NR | NR | Matheri et al., 2022
[5 empty] | NR | SAR-CoV-2, Hepatitis A and E, Polyomavirus, Papillomavirus, Enterovirus, Rotavirus, Adenovirus, Parechovirus, Salivirus, Astrovirus, and Zika | sewage water | 3.28 x 10^6 SARS-CoV-2 genome particles per liter of sample | NR | NR | Mainardi et al., 2022
[5 empty] | NR | RNA virus: SARS-CoV-2 and norovirus |  | NR | NR | NR
[5 empty] | NR | SARS-CoV-2, influenza, respiratory syncytial virus and measles RNA, pepper mild mottle virus (PMMoV) |  | High | High | NR | Hayes et al., 2024
|  | Multiplex PCR : Biofire® Filmarray® | 2 | 8 | costly | SARS-CoV-2 | nasal/ oropharyngeal | 93 | 100 | 3.30* 10^2 RNA copies/mL | Mahmoudi et al., 2022
|  | Droplet Digital PCR (ddPCR) | 2 | 7 | NR | norovirus, adenovirus, E. coli, ARGs |  | High | High | NR
|  | TaqMan® Pseudomonas aeruginosa Detection Kit: Pseudomonas CN | 1 | 7 | NR | P. aeruginosa in natural river |  | 93 |  | NR | Milligan et al., 2023
|  | Nucleic Acid Sequence-based Amplification-CRISPR Cleavage (NASBACC) | 1 | 7 | NR | ZIKV | plasma | NR | NR | 1x10^-15 mol/L (LOD) | Saltepe et al., 2022
| Isothermal Nucleic Acid Amplification (iNAD) | Loop-Mediated Isothermal Amplification (LAMP) | 6 | 7 | cost effective than PCR | Pseudomonas aeruginosa, Legionella pneumohila, Vibrio parahemolyticus, and Vibrio cholerae, ARGs |  | High | High | NR
[5 empty] | an IAM device developed for use in the field (e.g., using the rechargeable, battery operated AmpliFire®) currently cost ~$9000.00 | Legionella spp., Leptospira spp., etc. | wastewater | Good | Good | NR | Singh et al., 2024
[6 empty] | hepatitis c virus (HCV) | medical wastewater | High | High | NR | Chen et al., 2024
[6 empty] | Rhizopus | medical wastewater | NR | NR | 1 × 10^−1 pg/μL | Chen et al., 2024
[6 empty] | Cryptosporidium | medical wastewater | NR | NR | 1 × 10^−2~ 1 × 10^−3 pg/μL | Chen et al., 2024
[6 empty] | Tapeworm |  | NR | NR | 1 × 10^−2~ 1 × 10^−3 pg/μL | Chen et al., 2024
[6 empty] | SARS-Covid |  | NR | NR | 38 × 10−6 ng μL−1 | Donia et al., 2025
|  | RT-LAMP | 1 | 7 | cost effective than PCR | Avian Influenza Virus |  | High | High | NR
|  | Fluorescent In Situ Hybridization (FISH) | 2 | 7 | Very expensive | Salmonella spp., Enterobacteriaceae, E. coli, |  | High | High | NR | Singh et al., 2024
|  | Nucleic Acid Sequence-Based Amplification (NASBA) | 2 | 7 | cost effective than PCR |  |  | high (54.5%) | high (90.9%) | NR
|  | Recombinase polymerase amplification (RPA) | 1 | 5 | cost effective than PCR |  |  | high (62.5 %) | high (100%) | NR | Nieuwkerk et al., 2024
|  | Rolling circle amplification (RCA) | 1 | 3–5 | cost effective than PCR |  |  | high (57.1%) | high (100%) | NR | Nieuwkerk et al., 2024
|  | Helicase dependent amplification (HDA) | 1 | 3–5 | cost effective than PCR |  |  | high (100%) | high (100%) | NR
|  | MI-lF-RPA | 1 | 3–5 | NR | SARS-CoV-2 | Swab | 97 | 100 | 1 copy/µL | Mahmoudi et al., 2022
Sequencing | Metagenomics |  | 2 | 6 to 7 | NR | 43 different types of human viruses were identified in sewage sludge, including newly emerging viruses (Klassevirus, Coronavirus HKU1 and Cosavirus) associated with respiration and Enteroviruses |  |  |  | NR | Tiwari et al., 2024
| Next-Generation Sequencing (NGS) | Next-Generation Sequencing (NGS) | 3 | 7 | ~US $80 for 20 000 reads (amplicon sequencing) |  | Wastewater for raw sewage and final effluent | Very high | Very high | NR | Ofori et al; Bubba et al.
|  | Illumina and Oxford Nanopore Technologies (ONT) sequencing |  |  | NR | Antimicrobial-resistant Escherichia coli | Faecal sludge and Untreated wastewater |  |  | NR | Ofori et al., 2024
|  | Illumina DNA sequencing |  |  | NR | Antimicrobial resistance genes | Untreated wastewater | NR | NR | NR | Ofori et al., 2024
| Sanger Sequencing |  | 2 | 6 to 7 | very expensive |  |  | High | High | NR | Bubba 2023; Corrin et al., 2023
Immunoassay-Based Methods | ELISA (Enzyme-Linked Immunosorbent Assay) |  | 3 | 7 to 8 | NR | C. difficile | medical wastewater | 66-90% | NR | NR | Chen et al., 2024
| EIA (Enzyme Immunoassay) |  | 1 | 7 | NR | Clostridium difficile | medical wastewater | 50-90 % | NR | NR | Chen et al., 2024
| Immunofluorescence Assays (IFA) |  | 1 | 6 | NR | S. aureus | medical wastewater | NR | NR | 1 × 10^−2 CFU/L | Chen et al., 2024
[5 empty] | costly | human adenovirus; Giardia |  | NR | NR | NR
Biosensor-Based Detection | Electrochemical Biosensors | Electrochemical Biosensors | 6 | 6 | cost effective | bacteria (e.g., E. coli , Salmonella, Staphylococcus aureus), viruses (e.g., enterovirus, respiratory viruses, hepatitis virus), parasites, and fungi | medical wastewater | Fair. | Fair. | NR | Chen et al., 2024
[6 empty] | E Coli. | wastewater | NR | NR | 2.5 × 10^3 CFU/ml | Singh et al., 2025
[6 empty] | Shigella dysenteriae | wastewater | NR | NR | 1 CFU/ml | Singh et al., 2025
[6 empty] | SARS-CoV-4 | wastewater | NR | NR | 200 copies/ml (Calixarene functionalised graphene-based); I copy/ μl (Hybridization of rolling circle amplicons with a functionalized probe); 6.9 copies/μl (Antisense oligonucleotides) | Singh et al., 2025
|  | Electrochemical Impedance Spectroscopy (EIS) |  |  |  | E coli. |  | NR | NR | 10^3 CFU mL−1 (T2 phage); ∼14 CFU mL−1 (M13 phage) | Bayat et al., 2021
|  | Conductive biosensor |  |  |  | E. coli O157 | medical wastewater | NR | NR | 1 × 10^−2~1 × 10^−1 cells/L | Chen et al., 2024
[6 empty] | Bacillus anthracis | medical wastewater | NR | NR | 1 × 10^−3 CFU/L | Chen et al., 2024
|  | Electrochemical luminescent biosensor |  |  |  | Shiga toxin E. coli | medical wastewater | NR | NR | 0.3 pmol/L | Chen et al., 2024
[6 empty] | Mycotoxin | medical wastewater | NR | NR | 0.034 pg/mL | Chen et al., 2024
|  | Volt-ampere type biosensor |  |  |  | E. coli O157 | medical wastewater | NR | NR | 2.5 × 10^−2 CFU/L | Chen et al., 2024
[6 empty] | Schistosomiasis antigen | medical wastewater | NR | NR | 1.13 × 10^1~2.3 × 10^3 ng/mL | Chen et al., 2024
|  | Impedance |  |  |  | Escherichia coli |  | NR | NR | 2 × 10^4 CFU/ml | Singh et al., 2025
[6 empty] | Norovirus |  | NR | NR | 7.8 copies/ml | Singh et al., 2025
[6 empty] | Salmonella typhi | medical wastewater | NR | NR | 7.3 × 10^−2 CFU/L | Chen et al., 2024
[6 empty] | E. coli DH5a, Enterobacter cloacae, B. subtilis, Saccharomycetes | medical wastewater | high sensitivity | NR | NR | Chen et al., 2024
| Optical Biosensors | Optical Biosensors | 3 | 6 |  | Hepatitis A virus | wastewater | Fair | Fair | 1.1 pM | Singh et al., 2025
[6 empty] | Norovirus like particles | wastewater | NR | NR | 0.01 ng/ml | Singh et al., 2025
[6 empty] | Norovirus | wastewater | NR | NR | 12.1 × 10^-15 g/ml; 30 viruses/ml | Singh et al., 2025
|  | Fluorescent |  |  |  | Salmonella enteritidis | wastewater | NR | NR | 60 CFU/ml | Singh et al., 2025
[6 empty] | S. typhimurium | wastewater | NR | NR | 50 CFU/ml | Singh et al., 2025
[6 empty] | P. aeruginosa | wastewater | NR | NR | 1 CFU/ml | Singh et al., 2025
|  | photoelectrochemical |  |  |  | E. coli | medical wastewater | NR | NR | 4.6 × 10^1 CFU/L | Chen et al., 2024
[6 empty] | Vibrio parahemolyticus | medical wastewater | NR | NR | 2.5 × 10^1 CFU/L | Chen et al., 2024
[6 empty] | E. coli O157:H7 | wastewater | NR | NR | 1.125 CFU/ml | Singh et al., 2025
| Piezoelectric |  |  |  |  | Hepatitis A virus | wastewater | NR | NR | 1 × 10^5 viruses/ml | Singh et al., 2025
[6 empty] | Staphylococcus aureus | wastewater | NR | NR | 41 CFU/ml | Singh et al., 2025
| Phage-Based Biosensors: Use engineered bacteriophages for pathogen detection. | Electrochemical impedance spectroscopy (EIS); phage: M13 (wild-type) | 1 | 5 |  | E. coli XL1-Blue, E. coli K12 | Artificial river water | good | good | ∼14 CFU mL−1; stability: 16 days at 45 °C, very good to high stability in pH 3–10 | Bayat et al., 2021
|  | Electrochemical; phage: T7 reporter phage |  |  |  | E. coli (ECOR-13) | Drinking water |  |  | 105 CFU mL−1 within 4 hours, 1 CFU per 100 mL in 12 hours; The NRGp7-MBs were stable at 4 °C for 2 weeks | Bayat et al., 2021
|  | Colorimetric ; Engineered M13 |  |  |  | E. coli ER2738, P. aeruginosa, V. cholerae, Xanthomonas campestris | Seawater and tap water |  |  | 100 cells; NA | Bayat et al., 2021
|  | Colorimetric and luminescent assay; Engineered reporter T7 phage |  |  |  | E. coli BL21 | Drinking water |  |  | 1 CFU mL−1; NA | Bayat et al., 2021
|  | Colorimetric; Engineered reporter T7 phage |  |  |  | E. coli | LB media |  |  | 1.2 × 10^3 CFU mL−1 in 3.5 hours (without enrichment) 10^1 CFU mL−1 in 8 hours (with enrichment) | Bayat et al., 2021
|  | Bioluminescent assay; Engineered reporter phage T7 |  |  |  | E. coli BL21 | Lake water |  |  | <10 CFU mL-1 | Bayat et al., 2021
|  | Electrochemical Engineered T7 phages |  |  |  | E. coli BL22 | Drinking water |  |  | 10^5 CFU mL−1 in 3 hours 10^2 CFU mL−1 in 7 hours | Bayat et al., 2021
|  | Colorimetric Engineered T7 |  |  |  | E. coli BL23 | LB broth |  |  | 10 CFU mL−1 | Bayat et al., 2021
|  | Bioluminescent assay B. anthracis reporter phage (Wβ::luxAB-2) |  |  |  | Bacillus anthracis spores | Pond and lake water |  |  | 10 and 10^2 CFU mL−1 | Bayat et al., 2021
|  | Electrochemical impedance spectroscopy (EIS) T2 phage (wild type) |  |  |  | B strain of E. coli (ATCC 11303) | PBS buffer |  |  | 10^3 CFU mL−1 | Bayat et al., 2021
|  | Coliphage ECML-117 |  |  |  | E. coli O157: H7 | Municipal and agricultural water |  |  | 10^2 CFU mL−1 | Bayat et al., 2021
|  | Engineered T7 phage |  |  |  | E. coli BL21 | Broth and river water |  |  | 10^2–10^3 CFU mL−1 | Bayat et al., 2021
|  | Fluorescent Phage; HK620::PrrnB-gfp and P22::PrrnB-gfp |  |  |  | E. coli TD2158 and Salmonella enterica typhimurium 14028S | Sea water |  |  | 10 cells per mL | Bayat et al., 2021
|  | Colorimetric; T7 phage |  |  |  | E. coli BL21 | Drinking water |  |  | 10^4 CFU mL−1 and 10 CFU mL−1 (6 hour of preenrichment) | Bayat et al., 2021
|  | Fluorescent; Tetracysteine tagged M13 (M13KE-TC phage) |  |  |  | E. coli ER2738 | Drinking water |  |  | 1 CFU mL−1 in 40 mL of drinking water | Bayat et al., 2021
|  | Luminescent and colorimetric; Wild-type and engineered T4 |  |  |  | E. coli | Water |  |  | 40 or <10 CFU mL−1 | Bayat et al., 2021
|  | Bioluminescent; Engineered T7 |  |  |  | E. coli | Drinking water |  |  | <20 CFU mL−1 in 100 mL water | Bayat et al., 2021
|  | Colorimetric; Engineered M13KE phage |  |  |  | E. coli K12 (ER2378) | Drinking water |  |  | 50 CFU L−1 | Bayat et al., 2021
| Quartz Crystal Microbalance (QCM) |  | 1 | 3 |  | E. coli | Vermicompost tea, PBS | NR | NR | 10^2, 10^6 CFU mL^-1 (ligand- Antibody); 1.46* 10^3 CFU/mL (ligand- Aptamer) | Bonyadi et al., 2023
[6 empty] | A. actinomycetemcomitans | PBS |  |  | 800 CFU/ml | Bonyadi et al., 2023
[6 empty] | A. hydrophila | Fish tissue extract |  |  | 6 * 10^6 to 10^8 CFU | Bonyadi et al., 2023
[6 empty] | H3N2 canine influenza virus (CIV) | Diluted saliva |  |  | 2^2 HAU | Bonyadi et al., 2023
[6 empty] | S. typhimurium | PBS |  |  | 10^3 CFU/ml | Bonyadi et al., 2023
[6 empty] | L. innocua | PBS |  |  | 1.6*10^3 CFU/mL | Bonyadi et al., 2023
[6 empty] | Classical Swine fever Virus | PBS |  |  | 1.7 lg/mL | Bonyadi et al., 2023
[6 empty] | E. canis | PCR product |  |  | 22 copy number | Bonyadi et al., 2023
[6 empty] | B. cereus | PBS/ culture medium |  |  | NR | Bonyadi et al., 2023
| Colorimetric Biosensors |  | 3 | 6 |  | Pseudomonas aeruginosa |  | Medium | Medium | NR
[6 empty] | P. aeruginosa |  |  |  | NR
[6 empty] | E. coli ER2738, P. aeruginosa, V. cholerae, Xanthomonas campestris |  |  |  | NR | Bayat et al., 2021
[6 empty] | Pseudomonas aeruginosa | wastewater |  |  | 50 CFU/ml | Singh et al., 2025
Spectroscopy and Microscopy-Based Methods | Raman Spectroscopy | Single-Cell Raman Spectroscopy | 2 | 3 | set up cost high, individual analysis cost low | V. parahaemolyticus | wastewater | Medium | Medium | 10 CFU/ml | Singh et al., 2025
[6 empty] | Legionella spp., Bacillus spp., Pseudomonas spp., and Burkholderia spp |  |  |  | tree-based multiclass support vector machines (MCSVM) to discriminate Legionella at the species level with 87.3% accuracy; combining SCRS with SVM, the identification accuracies were up to 85% for P. aeruginosa at strain level from independent samples.
|  | Surface-Enhanced Raman Spectroscopy (SERS) | 2 | 3 |  | P. aeruginosa | wastewater | High | High | 20 CFU/ml | Singh et al., 2025
[6 empty] | S. enterica | wastewater |  |  | 1.5 CFU/ml | Singh et al., 2025
[6 empty] | Staphylococcus aureus | wastewater |  |  | 5 CFU/ml | Singh et al., 2025
|  | Raman Stable Isotope Probing | 2 | 3 |  |  | wastewater | High | High | NR | Singh et al., 2025
| Mass Spectrometry (e.g., MALDI-TOF MS) |  | 2 | 8 | very expensive | Aeromonas spp. | surface water and wastewater | 74-87.7 | Very high | NR | Milligan et al., 2023
[6 empty] | V. cholerae, V. alginolyticus, S. typhi. | wastewater | NR | NR |  | Singh et al., 2024
| Microscopy-Based Detection |  | 2 | 5 | NR | Ascaris spp. | medical wastewater | NR | Medium | 50 seeded eggs/gr faeces | Chen et al., 2024
Microfluidics | Paper-Based Devices (µPADs): Low-cost, disposable devices integrating CRISPR or LAMP. |  | 5 | 7 | ~7.5 USD (assay cost) | SARS-CoV-2 | Swab | 100 | 100 | 6.9 copies/mL(LOD) | Mahmoudi et al., 2022
| Lateral flow devices |  | 2 | 5 to 6 | NR | SARS-CoV-3 | Swab/ sputum | 87.69/ 91.94 | 95.01/ 92.41 | 500 copy/mL | Mahmoudi et al., 2022
Emerging and Integrated Methods | Lab-on-a-Chip (LOC) |  | 2 | 5 to 6 | NR | SARS-CoV-2 | Swab | 100 | 98.6 | NR | Mahmoudi et al., 2022
| Integrated RT-LAMP Microfluidic Chip |  | 1 | 6 | NR | SARS-CoV-2 | wastewater | 85.7 | High | NR | Mahmoudi et al., 2022
[5 empty] | ~2 USD (assay cost) | SARS-CoV-3 | swab | 100 | 100 | 100 GEa/mL | Mahmoudi et al., 2022
| CRISPR-Based Detection (e.g., SHERLOCK, DETECTR) |  | 3 | 3 | NR | ZIKV, DENV | serum, urine, body fluids | Very high | Very high | 2x10^-18 mol/L (LOD) (serum, urine); 8x10^-21 mol/L (LOD) (body fluids) | Saltepe et al., 2022
[5 empty] | costly | EBOV, LASV | blood, urine, saliva |  |  | 10 copies/μl | Saltepe et al., 2022
[5 empty] | NR | Plasmodium species | blood, human serum |  |  | 0.3 parasite/μl | Saltepe et al., 2022
[5 empty] | NR | SARS-CoV-2 | Nasopharyngeal swab |  |  | 42 copies/reaction | Saltepe et al., 2022
[5 empty] | NR | HPV16, HPV18 | anal swab | NR | NR | 1x10^-18 mol/L (LOD) | Saltepe et al., 2022
[5 empty] | NR | SARS-CoV-2 | Nasopharyngeal swab | NR | NR | 10 copies/μl | Saltepe et al., 2022
| Aptamer-Based Electrochemical Chips |  | 2 | 6 | NR | SARS-CoV-3 |  | High | High | NR | Donia et al., 2025
| Smartphone-Based Detection (Links biosensors to smartphones for real-time monitoring) |  | 2 | 6 | NR | SARS-CoV-2 | RT-PCR amplicons | Medium | Medium | 1 copy number/mL | Mahmoudi et al., 2022
| Xpert® Xpress |  | 1 | 6 | NR | SARS-CoV-3 | upper respiratory specimens | 99 | 97 | 0.0200 PFU/mL | Mahmoudi et al., 2022
| ID NOW |  | 1 | 6 | NR | SARS-CoV-2 | Nasal swab | 84 | 100 | 3900-20,000 gene copies/mL | Mahmoudi et al., 2022
| RTisochipTM-W |  | 1 | 6 | NR | SARS-CoV-2 | swab | NR | NR | 50 copies/mL | Mahmoudi et al., 2022
| one-Hour Low-cost Multipurpose highly Efficient System (HOLMES) |  | 1 | 6 | NR | PRV, JEV | synthetic nucleic acid, urine | NR | NR | 1x10^-17 mol/L | Saltepe et al., 2022
"NR" = not reported.

## References
WBE papers for review – detection methods
BAYAT, F., DIDAR, T. F. & HOSSEINIDOUST, Z. 2021. Emerging investigator series: bacteriophages as nano engineering tools for quality monitoring and pathogen detection in water and wastewater. Environmental Science-Nano, 8.
BHALLA, N., PAN, Y. W., YANG, Z. G. & PAYAM, A. F. 2020. Opportunities and Challenges for Biosensors and Nanoscale Analytical Tools for Pandemics: COVID-19. Acs Nano, 14, 7783-7807.
BONYADI, F., KAVRUK, M., UCAK, S., CETIN, B., BAYRAMOGLU, G., DURSUN, A. D., ARICA, Y. & OZALP, V. C. 2023. Real-Time Biosensing Bacteria and Virus with Quartz Crystal Microbalance: Recent Advances, Opportunities, and Challenges. Critical Reviews in Analytical Chemistry.
BUBBA, L., BENSCHOP, K. S. M., BLOMQVIST, S., DUIZER, E., MARTIN, J., SHAW, A. G., BAILLY, J. L., RASMUSSEN, L. D., BAICUS, A., FISCHER, T. K. & HARVALA, H. 2023. Wastewater Surveillance in Europe for Non-Polio Enteroviruses and Beyond. Microorganisms, 11.
CHEN, B. Y., HE, J. H., TIAN, K. W., QU, J., HONG, L. H., LIN, Q., YANG, K., MA, L. & XU, X. L. 2024. Research Progress on Detection of Pathogens in Medical Wastewater by Electrochemical Biosensors. Molecules, 29.
CORRIN, T., RABEENTHIRA, P., YOUNG, K. M., MATHIYALAGAN, G., BAUMEISTER, A., PUSSEGODA, K. & WADDELL, L. A. 2024. A scoping review of human pathogens detected in untreated human wastewater and sludge. Journal of Water and Health, 22, 436-449.
CUI, L., LI, H. Z., YANG, K., ZHU, L. J., XU, F. & ZHU, Y. G. 2021. Raman biosensor and molecular tools for integrated monitoring of pathogens and antimicrobial resistance in wastewater. Trac-Trends in Analytical Chemistry, 143.
DEVIANTO, L. A. & SANO, D. 2023. Systematic review and meta-analysis of human health-related protein markers for realizing real-time wastewater-based epidemiology. Science of the Total Environment, 897.
DONIA, A., MTHETHWA-HLONGWA, N. P., AMOAH, I. D., KUMARI, S. & BUX, F. 2024. Point-of-care diagnostics for SARS-CoV-2 wastewater-based epidemiology: a big leap toward miniaturization. Environmental Science-Water Research & Technology, 11, 10-28.
HAYES, E. K. & GAGNON, G. A. 2024. From capture to detection: A critical review of passive sampling techniques for pathogen surveillance in water and wastewater. Water Research, 261.
JULIAN, T. R., DEVAUX, A. J., BRÜLISAUER, L., CONFORTI, S., RUSCH, J. C., GAN, C. R., BAGUTTI, C., STADLER, T., KOHN, T. & ORT, C. 2024. Monitoring an Emergent Pathogen at Low Incidence in Wastewater Using qPCR: Mpox in Switzerland. Food and Environmental Virology, 16, 269-279.
KANTOR, R. S. & JIANG, M. X. 2024. Considerations and Opportunities for Probe Capture Enrichment Sequencing of Emerging Viruses from Wastewater. Environmental Science & Technology, 58, 8161-8168.
KARAMATI, E., LAW, I., WEESE, J. S., MCCARTHY, D. T. & MURPHY, H. M. 2024. Passive sampling of microbes in various water sources: A systematic review. Water Research, 266.
MAHMOUDI, T., NAGHDI, T., MORALES-NARVAEZ, E. & GOLMOHAMMADI, H. 2022. Toward smart diagnosis of pandemic infectious diseases using wastewater-based epidemiology. Trac-Trends in Analytical Chemistry, 153.
MAINARDI, P. H. & BIDOIA, E. D. 2022. Risks associated with the presence of SARS-CoV-2 in sewage and possible approaches to limit its spread through aquatic matrices. Vigilancia Sanitaria Em Debate-Sociedade Ciencia & Tecnologia, 10, 77-92.
MAO, K., ZHANG, H., PAN, Y. W. & YANG, Z. G. 2021. Biosensors for wastewater-based epidemiology for monitoring public health. Water Research, 191.
MATHERI, A. N., BELAID, M., NJENGA, C. K. & NGILA, J. C. 2023. Water and wastewater digital surveillance for monitoring and early detection of the COVID-19 hotspot: industry 4.0. International Journal of Environmental Science and Technology, 20, 1095-1112.
MILLIGAN, E. G., CALARCO, J., DAVIS, B. C., KEENUM, I. M., LIGUORI, K., PRUDEN, A. & HARWOOD, V. J. 2023. A Systematic Review of Culture-Based Methods for Monitoring Antibiotic-Resistant <i>Acinetobacter</i>, <i>Aeromonas</i>, and <i>Pseudomonas</i> as Environmentally Relevant Pathogens in Wastewater and Surface Water. Current Environmental Health Reports, 10, 154-171.
NIEUWKERK, D. M., KORAJKIC, A., VALDESPINO, E. L., HERRMANN, M. P. & HARWOOD, V. J. 2020. Critical review of methods for isothermal amplification of nucleic acids for environmental analysis. Journal of Microbiological Methods, 179.
OESCHGER, T. M., MCCLOSKEY, D. S., BUCHMANN, R. M., CHOUBAL, A. M., BOZA, J. M., MEHTA, S. & ERICKSON, D. 2021. Early Warning Diagnostics for Emerging Infectious Diseases in Developing into Late-Stage Pandemics. Accounts of Chemical Research, 54, 3656-3666.
OFORI, B., AGOHA, R. K., BOKOE, E. K., ARMAH, E. N. A., MORANGA, C. M. & SARPONG, K. A. N. 2024. Leveraging wastewater-based epidemiology to monitor the spread of neglected tropical diseases in African communities. Infectious Diseases, 56, 697-711.
RAMES, E., ROIKO, A., STRATTON, H. & MACDONALD, J. 2016. Technical aspects of using human adenovirus as a viral water quality indicator. Water Research, 96, 308-326.
SALTEPE, B., WANG, L. & WANG, B. J. 2022. Synthetic biology enables field-deployable biosensors for water contaminants. Trac-Trends in Analytical Chemistry, 146.
SANGKHAM, S. 2021. A review on detection of SARS-CoV-2 RNA in wastewater in light of the current knowledge of treatment process for removal of viral fragments. Journal of Environmental Management, 299.
SINGH, R., RYU, J., LEE, W. H., KANG, J. H., PARK, S. & KIM, K. 2025. Wastewater-borne viruses and bacteria, surveillance and biosensors at the interface of academia and field deployment. Critical Reviews in Biotechnology, 45, 413-433.
SINGH, S., AHMED, A. I., ALMANSOORI, S., ALAMERI, S., ADLAN, A., ODIVILAS, G., CHATTAWAY, M. A., SALEM, S. B., BRUDECKI, G. & ELAMIN, W. 2024. A narrative review of wastewater surveillance: pathogens of concern, applications, detection methods, and challenges. Frontiers in Public Health, 12.
SYMONDS, E. M. & BREITBART, M. 2015. Affordable Enteric Virus Detection Techniques Are Needed to Support Changing Paradigms in Water Quality Management. Clean-Soil Air Water, 43, 8-12.
TIWARI, A., KROLICKA, A., TRAN, T. T., RÄISÄNEN, K., ASMUNDSDÓTTIR, A. M., WIKMARK, O. G., LOOD, R. & PITKÄNEN, T. 2024. Antibiotic resistance monitoring in wastewater in the Nordic countries: A systematic review. Environmental Research, 246.
ZAHEDI, A., MONIS, P., DEERE, D. & RYAN, U. 2021. Wastewater-based epidemiology-surveillance and early detection of waterborne pathogens with a focus on SARS-CoV-2, <i>Cryptosporidium</i> and <i>Giardia</i>. Parasitology Research, 120, 4167-4188.
ZHANG, S. X., SHI, J. H., LI, X., TIWARI, A., GAO, S. H., ZHOU, X., SUN, X. Y., O'BRIEN, J. W., COIN, L., HAI, F. S. & JIANG, G. M. 2023. Wastewater-based epidemiology of Campylobacter spp.: A systematic review and meta-analysis of influent, effluent, and removal of wastewater treatment plants. Science of the Total Environment, 903.